## Description
Registered Compositing Planner Scenario
Instructions for use test scenarios
Module Overview
Description:
This test scenario provides the registered compositing planner actor with RO-DC1, MMRO-II-2 and RAD 4.8 and receives RO-DC2 and RO-DC3
- Checks the correct storage of Single Series CT dataset (DICOM association & interaction level)
- Checks the correct storage of Spatial Registration dataset (DICOM association & interaction level)
- Checks the correct storage of General Dose dataset (DICOM association & interaction level)
- Validation of the created Single Plan Dose (RO-DC3) (including RO critical modules)
- Validation of the created Composite Dose (RO-DC2) (including RO critical modules)
- Verifies the critical attribute mapping requirements
When this scenario starts the following dialog will show up.
Select the directory where the images are located and where the storage files need to be placed, for the other IOD's a file needs to be selected.
Transactions sent to actor:
MMRO 4 (Registered Structure Set Retrieval)
RO-DC1 (General Dose Retrieval)
MMRO 2 (Spatial Registration Retrieval)
RAD-4.8 (Modality Images Stored)
Expected transactions from actor:
RO-DC3 (Single Plan Dose Storage)
RO-DC2 (Composite Dose Storge)
Used scenario dataset:
This scenario uses the selected test dataset.

## Module Overview
Registered Compositing Planner Scenario Rule Overview
RT Dose
C.7.1.1 Patient Module
C.7.2.1 General Study Module
C.8.8.1 RT Series Module
C.7.4.1 Frame Of Reference Module
C.7.5.1 General Equipment Module
C.7.6.1 General Image Module
C.7.6.2 Image Plane Module
C.7.6.3 Image Pixel Module
C.7.6.6 Multi-Frame Module
C.8.8.3 RT Dose Module
Attribute | Tag |  | Description | Rules
Dose Type | 3004,0004 | R+* | If any contributing dose has Dose Type = EFFECTIVE, the composite dose must have Dose Type = EFFECTIVE | DoseTypeRule, ValueListRule
Dose Summation Type | 3004,000A | R+* | Shall have the value PLAN or MULTI_PLAN | ValueListRule
Tissue Heterogeneity Correction | 3004,0014 | R+ | Shall be present but may be null. The value shall be given if known. | TissueHeterogeneityCorrectionRule
C.12.1 SOP Common Module

## General Modules
Rule overview
C.7.1.1 Patient Module
Attribute | Tag | Type | Description | rule name
Patient's Name | 0010,0010 | R+ | Same as in reference data | CriticalAttributeMappingRule
Patient ID | 0010,0020 | R+ | Same as in reference data | CriticalAttributeMappingRule
Patient's Birthdate | 0010,0030 | O+ | Same as in reference data | CriticalAttributeMappingRule
Patient's Sex | 0010,0040 | O+ | Same as in reference data | CriticalAttributeMappingRule
C.7.2.1 General Study Module
Attribute | Tag |  | Description | rule name
Study Instance UID | 0020,000D | R+* | Same as in reference data | CriticalAttributeMappingRule
Study Date | 0008,0020 | R+ | Same as in reference data | CriticalAttributeMappingRule
Study Time | 0008,0030 | R+ | Same as in reference data | CriticalAttributeMappingRule
'Referring(Physician) | 0008,0090 | O+ | Same as in reference data | CriticalAttributeMappingRule
Study ID | 0020,0010 | R+ | Same as in reference data | CriticalAttributeMappingRule
Accession Number | 0008,0050 | O+ | Same as in reference data | CriticalAttributeMappingRule
Study Description | 0008,1030 | O+ | Same as in reference data | CriticalAttributeMappingRule
C.7.4.1 Frame Of Reference Module
Attribute | Tag |  | Description | rule name
Frame of Reference UID | 0020,0052 | R+* | The Registration instance will have the same Frame of Reference UID as the registered image set. | IsRequiredRule, HasValueRule
Position Reference Indicator | 0020,1040 | O+* | The same comments as the Frame of Reference UID apply here. | HasValueIfPresentRule
C.7.5.1 General Equipment Module
Attribute | Tag |  | Description | rule name
Manufacturer | 0008,0070 | R+ | IHE requires that this element be present, and should contain the manufacturer of the equipment creating the structure set, plan, or dose. If the equipment is storing and forwarding information, the value of this element shall be preserved. If a new plan is created from a previous plan, the manufacturer of the equipment producing the new plan shall insert their identifier in this element. If a new structure set is created from a previous structure set, the manufacturer of the equipment producing the new structure set shall insert their identifier in this element. | ManufacturerMappingRule
Manufacturer's Model Name | 0008,1090 | R+ | If an application resamples and re-exports a series of CT images, or modifies an instance then this element must be present, and must contain the model name of the equipment doing the resampling. | ManufacturerMappingRule
C.7.6.1 General Image Module
Attribute | Tag |  | Description | rule name
InstanceNumber | 0020,0013 | R+ |  | IsRequiredRule
C.7.6.2 RT Image Plane Module
Attribute | Tag |  | Description | rule name
Image Orientation (Patient) | 0020,0037 | R+* | 1. Shall be present in every RT Dose IOD. 2. For an axial image (Image Type, (0008,0008) should be AXIAL), direction cosines shall be (+/-1, 0, 0, 0, +/1, 0) with an angle tolerance of 0.001 radians (0.057 degrees) | ImageOrientationPatientRule
Slice Thickness | 0018,0050 | O+* | Shall not be relied on. | HasValueIfPresentRule
Slice Location | 0020,1041 | O+* | Shall not be relied on. | HasValueIfPresentRule
Pixel Spacing | 0028,0030 | R+* |  | IsRequiredRule
C.7.6.3 Image Pixel Module
Attribute | Tag |  | Description | rule name
SamplesperPixel | 0028,0002 | R+* |  | IsRequiredRule
C.7.6.6 Multi-Frame Module
Attribute | Tag |  | Description | rule name
Frame Increment Pointer | 0028,0009 | R+* |  | IsRequiredRule, FrameIncrementPointerRule
C.8.8.3 RT Dose Module
Attribute | Tag |  | Description | rule name
Samples per Pixel | 0028,0002 | R+* | Value shall be 1 | ValueRule
Photometric Interpretation | 0028,0004 | R+* | Value shall be MONOCHROME2 | ValueRule
Bits Allocated | 0028,0100 | R+* | Value should be 16 or 32 | ValueListRule
Bits Stored | 0028,0101 | R+* | Value equal to Bits Allocated (0028,0100) | EqualToRule
High Bit | 0028,0102 | R+* | Value = Bits Stored(0028,0101) - 1 | HighBitRule
Pixel Representation | 0028,0103 | R+* | Shall have the value 0 = unsigned integer. Negative dose values shall not be present." | ValueRule
Dose Units | 3004,0002 | R+* | Shall be equal to the enumerated value GY | ValueRule
Dose Comment | 3004,0006 | R+ | "Shall be present if Referenced RT Plan Sequence is NOT PRESENT If present it should have the same value as RT Plan Description and it should not be empty" | DoseCommentRule
Normalization Point | 3004,0008 | O+* | Shall not be relied on. | HasValueIfPresentRule
Referenced RT Plan Sequence | 300C,0002 | R+* | Shall be present if Dose Summation Type (3004,000A) has the value PLAN | IsRequiredRule
">>Referenced Fraction Group Sequence | 300C,0020 | R+* | Shall be present if the parent sequence is present, and shall reference a single fraction group within the referenced RT Plan. Referenced Frame of Reference | ReferencedFractionGroupSequenceRule
Grid Frame Offset Vector | 3004,000C | R+* | 1. First z coordinate shall be equal to zero. 2. The remaining z coordinates shall be relative to the starting z position in Image Position (Patient) (0020,0032). Note: Can not be checked, the only thing that can be checked is that the offset is increasing every step. | GridFrameOffsetVectorRule
C.12.1 SOP Common
Attribute | Tag |  | Description | rule name
SOP Instance UID | 0008,0018 | R+* | If an application alters an Information Object instance, then the new Information Object instance shall be assigned a new UID. | IsRequiredRule, HasValueRule
Specific Character Set | 0008,0005 | O+* | Actors must not rely on the presence of this attribute to operate correctly. | HasValueIfPresentRule
Instance Creation Date | 0008,0012 | O+* | Actors must not rely on the presence of this attribute to operate correctly. | HasValueIfPresentRule
Instance Creation Time | 0008,0013 | O+* | Actors must not rely on the presence of this attribute to operate correctly. | HasValueIfPresentRule
Instance Creator UID | 0008,0014 | O+* | Actors must not rely on the presence of this attribute to operate correctly. | HasValueIfPresentRule
Instance Number | 0020,0013 | O+* | Actors must not rely on the presence of this attribute to operate correctly. | HasValueIfPresentRule
Rule overview