## BRTO - ROI Contour Module table
ROI Contour Module
Attribute Name | Tag | Presence | Attribute Note
ROI Contour Sequence | (3006,0039) | R* | -
>ROI Display Color | (3006,002A) | O+* | Not required - no compliant implementation shall rely on this element being present for proper operation. However applications are allowed to be aware of this element and use it to map display colors.
>Contour Sequence | (3006,0040) | R+* | Shall be present. Shall contain an item for each contour in the ROI. Compliant implementations shall be able to handle as many as 100 contours on a single slice. That is, the number of contours in items in all Contour Sequences with the same z-coordinate (and referenced CT image) should be less than or equal to 100.
>>Contour Image Sequence | (3006,0016) | R+* | Shall be present with a single item. This item is the image upon which this contour should be placed. If the contour type is CLOSED_PLANAR, then the z-coordinates of the contour shall match the z-coordinate of Image Position (Patient) in the image.
>>>Referenced SOP Class UID | (0008,1150) | R+* | Shall be present with a value of '1.2.840.10008.5.1.4.1.1.2'
>>>Referenced SOP Instance UID | (0008,1155) | R* | SOP Instance UID of the image being referenced.
>>>Referenced Frame Number | (0008,1160) | O+* | Shall not be present
>>Contour Geometric Type | (3006,0042) | R+* | Shall be present, with a value of POINT or CLOSED_PLANAR. Conforming implementations must properly interpret this value.
>>Contour Slab Thickness | (3006,0044) | O+* | Not required - no compliant implementation shall rely on this element being present for proper operation.
>>Contour Offset Vector | (3006,0045) | O+* | The profile requires that this attribute be zero if present.
>>Number of Contour Points | (3006,0046) | R+* | Required, and must match the actual number of points in Contour Data. Shall not exceed the number for which the ontour Data can not be encoded when using explicit transfer syntax.
>>Contour Data | (3006,0050) | R+* | Shall be present. If contour type is CLOSED_PLANAR, then all points must have the same z-coordinate. This z-coordinate shall match the z-coordinate in the related CT image within 0.01 mm (contained in the Contour Image sequence in the same item of the ROI Contour Sequence as this data). An implication of this is that the CLOSED_PLANAR contours are axial.